## Slide 1: Gas Chromatography

## Slide 2: Contents
Introduction
Hardware
GC injection
GC Basic Components
Applications

## Slide 3: Gas chromatograph
An analysis technique in which the sample mixture is injected, vaporized into a stream of carrier gas (as nitrogen or helium) moving through a column containing a stationary phase, separated into its component according to their affinity for the stationary phase and detected.

## Slide 4: Separations Process
Boiling points is the number one factor to consider in separation of compounds on the GCs.
Differences in polarity of the compounds is only important if you are separating a mixture of compounds which have widely different polarities.
Column temperature, the polarity of the column, flow rate, and length of a column are constant .

## Slide 5: Instrumentation
Manual Injection.
Automatic Injections.
Injectors
Columns
Oven
Detectors
Carrier Gas

## Slide 6: Samples injection system
Sample introduction by syringe
Most commonly used technique
Different syringe types manual and automatic
Other techniques and devices
Sampling valve (gas or liquids)
Head-space (liquids or solids)

## Slide 7: Manual Injection
Micro Syringes
Are used to introduce a known volume of a liquid or gas samples.
Adaptor
Can be used to help control the volume injected.
Syringe injection
Samples should be injected as a plug.
Rapid and consistent injection is necessary in order to obtain Acceptable precision
Injection Volume
Liquids 0.1-10 µl is typical
Gases 0.5- 5 ml is typical

## Slide 8: Hot Needle injection
Draw sample into syringe barrel.
Draw 2-3 ul air into barrel.
Inset needle into injection port and allow to heat
for a few seconds.
Rapidly inject sample and withdraw the needle.
This insure that all sample is injected and the hot
needle assists in solvent volatilization.

## Slide 9: Normal injection
Rinse the syringe with your sample at
least twice .
Draw up the suggested amount of sample into the syringe.
Pull up about 1-2 ul of air, This will give a signal showing the beginning of the elution.
Insert the needle through the injection
port and septum in one movement.
Quickly push the plunger.

## Slide 10: Auto Sampler
Large sample vial capacity
Max 2 trays installed
Simultaneously
1, 2 and 2.5 ml vials (up to 300 vials)
Syringe size:
5, 10, 100 and 250 µl.
Self recognized
syringes and trays
Washing station:
multiple Solvent rinsing supported
Rapid Mode: Allows to perform cleaning operations during GC run or cooling time.
Suitable for Ultra Fast GC requirements, eliminating dead times between successive runs.

## Slide 11: GC Basic Components
Injector
Column
Oven
Detector
Carrier
Gas

## Slide 12: Carrier gas
An inert gas, which is used to sweep a mixture to be separated through a gas chromatograph , (helium, hydrogen, or nitrogen).
Push the sample through the gas chromatograph column.
Clean out the gas chromatograph column after sample analysis.

## Slide 13: Carrier Gas Control
The Flow mode has four options for the carrier gas control:
Constant flow
Constant pressure
Programmed flow
Programmed pressure
The electronic control of the carrier gas allows also the following operations.
Column Evaluation
Gas Saver Function
Leak Check

## Slide 14: Gas Purity
All pure gases are classified by grade, so you can be certain of purity levels.
The first digit of the classification indicates the number of nines purity (for example, 5.0 = 99.999% purity).
The second digit is the number following the last nine (for example 4.7 helium has a guaranteed minimum purity of 99.997% and a corresponding maximum impurity level of 0.003% or 30ppm).

## Slide 15: Carrier Gas
The carrier gas
Ultra-pure and research-grade gases of up to 99.9999% (Grade 6.0) purity.
The carrier gas system often contains a molecular sieve to remove water or other impurities.
Linear Velocity (u)
Is the speed at which the carrier gas or mobile phase travels through the column.
The linear velocity is generally expressed in cm/s.
The linear velocity is independent of the column diameter while the flow rate is dependent on the column diameter.

## Slide 16: Required Gases Purities
Helium For carrier gas: 99.995%1 high purity, with less than 1.0 ppm each of water, oxygen, and total hydrocarbons after purification.
Use water, oxygen, and hydrocarbon traps.
Hydrogen For carrier or detector fuel gas: 99.995%1 high purity, with <
1.0 ppm of total hydrocarbons after purification.
Use water, oxygen and hydrocarbon traps.
Air For detector fuel gas: 99.995%1 high purity.
Air compressors are not acceptable because they do not
meet pressure, water, and hydrocarbon requirements.
Nitrogen For carrier or make-up gas: 99.995% high purity, with less than 1.0 ppm of total hydrocarbons after purification.
Argon 5% Methane For ECD make-up gas: 99.995%1 high purity.

## Slide 17: Leak Detection
CAUTION
Do not use liquid soap leak detectors to check for leaks. Liquid soap leak detectors may contaminate you system.
A mixture of 50% H2O/50% methanol or isopropyl
Alcohol may be used as a liquid leak detector.
WARNING! Never use liquid leak detectors on or around electronic pneumatic circuits

## Slide 18: Generators
Nitrogen Generator
The nitrogen generator can also operate directly from the laboratory compressed air supply.
General contaminants are first removed with appropriate filters and adsorbents and the purified air passes over layers of polymeric hollow fiber membranes through which nitrogen selectively permeates.
Hydrogen Generator
In the Packard Hydrogen Generator, hydrogen is generated electrolytically from pure deionized water.
The electrolysis unit uses a solid polymer electrolyte and thus does not need to be supplied with electrolytes, only the deionized water.

## Slide 19: Injector
The injection port Is a hollow, heated, glass-lined cylinder
The injector is heated so that all components in the
sample will be vaporized.
If the temperature is too low, separation is poor and broad spectral peaks should result or no peak develops at all.
If the injection temperature is too high, the specimen may decompose or change its structure.
The temperature of the sample port is usually about 50°C higher than the boiling point of the least volatile component of the sample.

## Slide 20: Injector

## Slide 21: Injection Techniques
Vaporizing.
The liquid sample is evaporated prior to be transferred to the separation column.
Split Spliless: SSL (permanently hot)
Programmed Temperature Vaporizer: PTV
Direct: PKD, PPKD (permanently hot, low resolution columns)
Nonvaporizing.
The liquid sample evaporates into the separation column (or a precolumn)
Cold On Column: OC (permanently cool)

## Slide 22: Septa
Ensure optimal performance of your GC instrument with bleed and temperature.
Made of low-bleed silicone, have excellent mechanical properties, are ideal for demanding GC and GC-MS applications, and may be used reliably up to 400 °C.
Septum must be replaced at least after 200 injections.

## Slide 23: Injector Types
Split/ Splitless Injector
On-Column Injector
High Oven Temperature On-Column Injector
Large Volume On-Column Injector
Packed Column Injector
Purged Packed Injector
Programmable Temperature Vaporizing Injector

## Slide 24: Split/Splitless Injector (vaporizing injector)
Split mode
The split vent is open, part of the sample go into the column.
When analyzing high concentration or neat samples.
Yields the sharpest peaks if the split gas is properly mixed.
Standard for capillary columns.

## Slide 25: Good injector
Capable to quantitative accept a broad volatility range.
Low discrimination.
To handle dirty and clean matrices.
With dirty matrices reduces sample clean-up and preserves the column.
With clean matrices boosts sensitivity with LVI techniques .
Extremely inert.
Able to handle polar/active compounds.
Provide optimum band shape.

## Slide 26: Column
The column
Is where the chromatographic separation of the sample occurs.
Several types of columns are available for different chromatographic applications:
The heart of the system.
It is coated with a stationary phase which greatly influences the separation of the compounds.

## Slide 27: Stationary phase
Solid resin packed in a column , or
Liquid supported by course paper or Inactive solid over which a mixture passes.
Each component of the mixture differs in the way it adheres to this phase and therefore travels along it at a unique rate.

## Slide 28: Adsorbents
There are two types of packing employed in GC, the adsorbents and the supports, on which the stationary phase is coated.
There are both inorganic and organic types of adsorbents.
Alumina, in an activated form, is used to separate the permanent gases and hydrocarbons up to about pentane.
Silica gel It is used for the separation of the lower molecular weight gases and some of the smaller hydrocarbons sulfur gases, hydrogen sulfide, sulfur dioxide and carbon disulfide.
Synthetic zeolites used for the separation of hydrogen, oxygen, nitrogen, methane and carbon monoxide and also rare gasses.

## Slide 29: Stationary Phase
POLYSILOXANES
The most common stationary phases.
They are available in the greatest variety and are the most stable, robust and versatile.
The most basic Poly Siloxane is the 100% methyl substituted
POLYETHYLENE GLYCOLS
They are less stable, less robust and have lower temperature limits than most Poly Siloxane.
must be liquids under GC temperature conditions.

## Slide 30: Inactive Solid Supports
There have been a number of materials used as supports for packed GC columns including,
Celite (a proprietary form of a diatomaceous earth), fire-brick (calcined Celite), fire-brick coated with metallic silver or gold, glass beads, Teflon chips and polymer beads.
Polystyrene beads

## Slide 31: Column Types
Conventional
1/8-1/4 OD
6-8 feet in length
Stainless steel or glass tube
Preparative
>1/4 OD
> 10 feet in length
Capillary
0.1- 0.5 ID
10 – 100 meters in length

## Slide 32: Factors Affecting Column Separations
Volatility of compound: Low boiling (volatile) components will travel faster through the column than will high boiling components
Polarity of compounds: Polar compounds will move more slowly, especially if the column is polar.
Column temperature: Raising the column temperature speeds up All the compounds in a mixture, “Columns have lower and upper temperature limits”.
Column packing polarity: Usually, all compounds will move slower on polar columns, but polar compounds will show a larger effect.
Flow rate of the gas through the column: Speeding up the carrier gas flow increases the speed with which all compounds move through the column.
Length of the column: The longer the column, the longer it will take all compounds to elute. Longer columns are employed to obtain better separation.

## Slide 33: Loss of Separation or Resolution
Contaminated column.
Damaged stationary phase.
Different column temperature, carrier
flow rate or column.
Large changes in the sample concentration.
Improper injector operation.

## Slide 34: Column Damage
Column breakage
Column bleed
Thermal damage
Oxygen damage
Chemical damage

## Slide 35: Oven
The use of a temperature programmed for
the column oven influences the separation process significantly and is used for optimization of time and peak separation.
The oven must not be opened when the oven temperature is above room temperature.
Never turn off the nitrogen flow unless the column and oven are at room temperature.

## Slide 36: The Oven Capabilities
Temperature range 5 °C above ambient to 350 °C
Temperature programming - up to six ramps
Maximum run time - 999.99 minutes
Temperature ramp rates - 0 to 120°C/min
The oven accommodates one inlet, one
detector, and one column

## Slide 37: Multiple-ramp temperature programs
A multiple-ramp temperature program changes the oven temperature from an initial value to a final temperature, but with various rates, times , and temperatures in between.
Multiple ramps can be programmed for temperature decreases as well as increases.

## Slide 38: Detector
The part of a gas chromatograph which signals the change in composition of the mixture passing through it.

## Slide 39: Detectors
Electron Capture Detector.
Flame ionization Detector.
Nitrogen Phosphors Detector.
Thermal Conductivity Detector.
Flame Photometric Detector.
Photo ionization Detector.
Electrolytic Conductivity Detector.
Mass Spectrometric Detector.

## Slide 40: Electron Capture Detector (ECD)
Mechanism:
Electrons are supplied from a 63Ni foil lining the detector cell. A current is generated in the cell. Electronegative compounds capture electrons resulting in a reduction in the current. The amount of current loss is indirectly measured and a signal is generated.
Selectivity:
Halogens, nitrates, conjugated carbonyls
Sensitivity:
0.1-10 pg (halogenated compounds);
1-100 pg (nitrates); 0.1-1 ng (carbonyls)
Nitrogen or argon/methane
300-400°C
Linear range: Gases: Temperature:

## Slide 41: Flame ionization Detector (FID)
Mechanism:
Compounds are burned in a hydrogen-air flame.
Carbon containing compounds produce ions that are attracted to the collector.
The No. of ions hitting the collector is measured and a signal is generated.
Selectivity:
Compounds with C-H bonds.
Sensitivity: Linear range: Gases: Temperature:
0.1-10 ng
105 -107
Combustion hydrogen and air; Makeup He or N2
250-300°C,and 400-450°C for high temp.

## Slide 42: Nitrogen Phosphors Detector (NPD)
Mechanism:
Compounds are burned in a plasma surrounding a rubidium bead supplied with hydrogen and air.
Nitrogen and phosphorous containing compounds produce ions that are attracted to the collector.
The number of ions hitting the collector is measured
and a signal is generated.
Selectivity: Nitrogen and phosphorous
Sensitivity: 1-10 pg
Linear range: 104 -106
Gases: Combustion - hydrogen and air; Makeup - Helium Temperature: 250-300°C

## Slide 43: Thermal Conductivity Detector (TCD)
Mechanism:
A detector cell contains a heated filament with an applied current. As carrier gas containing solutes passes through the cell, a change in the filament current occurs.
The current change is compared against the current in a reference cell.
The difference is measured and a signal is generated.
Selectivity:
All compounds except for the carrier gas
Sensitivity:
Linear range:
Gases:
Temperature:
5-20 ng
105 -106
Makeup - same as the carrier gas 150-250°C

## Slide 44: Detector
High sensitivity.
Rapidly respond to concentration changes.
Large linear range.
Stable with respect to noise and drift.
Low sensitivity to variation in flow, pressure and temperature.
Possible selectivity.
Produces an easily handled signal.
A temperature range from room temperature to at least 400 C.

## Slide 46: Major Current GC Markets

## Slide 47: Environmental
What is the environmental market?
Testing or commercial laboratories
Industrial laboratories
Government laboratories
Research institutes
Requirements of the applications
Drinking water
High sensitivity Waste
Sensitivity and selectivity Air
Sample introduction

## Slide 48: Clean water analysis
Pollutants in water
Halocarbons
Acid priority pollutants: phenols, chlorophenols, nitrophenols
Pesticides and PCBs
Base neutral priority pollutants
Polynuclear aromatic hydrocarbons

## Slide 49: Petrochemical and Gas
Large replacement business
Refinery
Oil Industry
Gas suppliers
Requirements of the applications
Multi-valves applications
Fastest possible cycle time
QC of gases: sensitivity
Customized software
Easy to use data handling and reporting

## Slide 50: Food & Beverages
Very extended market field
No really regulated methods
Ideal market to exploit the TRACE GC modularity
Requirements of the applications
QC of producers: Sensitivity and rapidity
Multi detection capability
Correct sample inlet (PTV-OC)
Easy to use data handling and reporting

## Slide 51: Pharmaceutical
Highly regulated market (pharmacopeia)
Requires Validation package
Requirements of the applications
Limited instrument requirements SSL/NPD/FID
Highly Automated market
High sensitivity detectors
A unified chromatography software